## Scientific equipment
Information on significant scientific equipment/devices and special premises of the Scientific-Research Institute as of January 1, 2022
Scientific Research Institute / Center | Significant scientific equipment/devices | [7 empty] | Special scientific fucilities/territories etc. (Tunnel / Experimental Field / Plot / Polygon)
Name | Title of the equipment/device (Manufacturer / Brand) | Technical characteristics | Date of issue | Exploitation staring year | Usage/application (works that can be performed with the device) | Purpose of usage/application (scientific / educational / industrial) | Technical condition (Operational - working condition / Zero usage / Requires repairment / Depreciated | Structure / Department / Laboratory responsible for the operation | Description (no more than 700 words)
E. Nanobashvili Laboratory of radiation chemistry | [8 empty] | The Radiation Chemistry Laboratory of the Institute of Inorganic Chemistry and Electrochemistry is located on the territory of the former nuclear reactor, which is surrounded by a brick fence. The facility is guarded by an armed state security service, entry is through a checkpoint. The laboratory is directly located in a special capital building designed for high-energy sources, only one floor of which is above ground (area - 574 m2), where laboratory rooms (No. 2-15) and a corridor are located. And two more floors are underground. This arrangement makes it possible to use the soil layer above the radiation sources as an additional biological protection. The design of the building also provides for the possibility of convenient use of lifting mechanisms in cases of installation or dismantling of important equipment in terms of dimensions and weight.On the -2nd floor, in special rooms (7 meters deep with a 3-meter concrete roof and the same walls), the gamma-device K-60000 and the electronic accelerator ЕLU-4 (currently not functioning) are installed. The facilities can be accessed through independent doors and a labyrinth. Equipment control panels are located on the ground floor of the building in room No. 1
| Gamma device K-60000; Manufacturer - Postbox /4110, Moscow, USSR | Co 60 source activity: - 50-60 curies (initial - 49,000 curies). | 1963 | 1964 | Gamma-ray radiation effects on various organic and inorganic systems; | Scientific | Working condition | E. Nanobashvili Laboratory of radiation chemistry
| Electron linear accelerator ЕLU-4; Manufacturer - city of Kostrom | Electron energy range: 3 - 5 MeV; Impulse current power of electrons: 1 ampere; Pulse current duration: 5 micro second. | 1987 | 1987 | Impact of high energy electrons on various organic and inorganic systems; | Scientific | Needs repair | E. Nanobashvili Laboratory of radiation chemistry
| Electrochemical Analyzer, CH Instrument (CHI 600E) | The potential control range is ±10 V and the current range is ±250 mA. The instrument is capable of measuring current down to picoamperes. The system contains a fast digital function generator, a direct digital synthesizer for high frequency AC waveforms, high speed dual-channel data acquisition circuitry and a potentiostat | 2016 | 2016 | For general purpose electrochemical measurements. Cyclic Voltammetry, Linear Sweep Voltammetry, Chronoamperometry, Chronocoulometry, Bulk Electrolysis with Coulometry, Open Circuit Potential – Time | Scientific | Active use | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Potentiostat/Galvanostat, CH Instrument (CHI 1140C) | The potential control range is ±10 V, the current range is ±2 A, and the compliance voltage is ±25 V. The system contains a high speed digital function generator, a fast data acquisition system, current signal filters, iR compensation circuitry, a potentiostat, and a galvanostat | 2019 | 2019 | The Model power potentiostat/galvanostat is designed for electrochemical applications that require relatively large current and high compliance voltage, such as battery studies, corrosion, electrolysis, and electroplating. Cyclic Voltammetry, Linear Sweep Voltammetry, Staircase Voltammetry, Tafel Plot, Chronoamperometry, Chronocoulometry, Differential Pulse Voltammetry, Normal Pulse Voltammetry, Square Wave Voltammetry, Amperometric i-t Curve, Double Differential Pulse Amperometry, Bulk Electrolysis with Coulometry. | Scientific | Active use | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Potentiostat, CH Instrument (CHI 1100C) | The potential control range is ±10 V, the current range is ±2 A, and the compliance voltage is ±25 V. The system contains a high speed digital function generator, a fast data acquisition system, current signal filters, iR compensation circuitry, a potentiostat. | 2019 | 2019 | The Model power potentiostat/galvanostat is designed for electrochemical applications that require relatively large current and high compliance voltage, such as battery studies, corrosion, electrolysis, and electroplating. Cyclic Voltammetry, Linear Sweep Voltammetry, Chronoamperometry, Chronocoulometry, Open Circuit Potential – Time, Bulk Electrolysis with Coulometry | Scientific | Active use | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Potentiostat pi-50-t with programmer, USSR | Maximum voltage of the potentiostat in static mode ± 50B; Maximum output current in static mode ± 1A; Maximum output current in potentiostat mode (pulse-not more than 300 ms frequency 50 Hz) ± 6A; Output current range 1 Ma (upper limit), 10 A, 0.1 A, 1A and 10A (last only pulse mode) | 1990 | 1991 | The purpose of the Potentiostat pi-50-t is to study fast electrochemical processes, at the electrode-electrolyte interface by potentiostatic, potentiodynamic, and impulse methods | Scientific | Active use | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Polarograph analyzer MA, Pzadelkis Budapest | standard potential ± 0.05-2.0V, diffusion compensation current ± 0.5-11a. Sensitive current - (1,0-3,5) \ 3.1010, (1.0-3.5) \ 3.10-6 | 1981 | 1982 | Polarograph analyzer used in potentiostatic and galvanostatic measurements of electrochemical processes | Scientific, Educational, Industrial | Depreciated, not subject to repair | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Press mechanism AURION, HOT SURFACES | Size 38 \ 38 cm, power 1300 W, temperature 0-2300C | 2011 | 2012 | Press mechanism creates pressure to compress the substance, A special material to change the shape of a substance. Can deplete fluid. It can create compression conditions of up to several kilograms. | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Potentiostat Ezstat Ezstat pro, NuVant Systems Inc | Power interval ± 1A; 10mA, 100mc, 1mc control potential ± 10A; Complex Volt ± 25 V, Step 40 PC A | 2011 | 2012 | Used for electrochemical measurements, for the study of corrosion processes, for the study of batteries | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Spectrometer DR- / 4000V, LASA HACH USA | Spectral Coefficient, Optical Density Range-B, Wavelength Range Nm | 2012 | 2013 | Spectrophotometer used in chemical laboratories for the determination of chemical elements in solutions and solid samples | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Photoelectrocalorimeter KFK-2 Concentration, USSR | 220 V, 50/60 Hz | 2006 | 2007 | Photoelectrocalorimeter is used to determine the concentration of a substance. Colorimeter output signal with all filters 15.0-105.0 | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Water purification system with accessories DURA series, HNiteche Pure Water technologies China | Production (250C) -12/24 l / h, main capacity water capacity 50 l. Electronic characteristics AC200-240B, 50-60 Hz | 2014 | 2015 | Water treatment system can produce 2 types of water: deionized and purified water | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Mechanical stirrer, Russian Federation | Amplitude 20 mm, Engine power 16 W, Engine speed limit 100-150 min | 2015 | 2016 | Mixing of laboratory solutions | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Geological jaw crusher, Russian Federation | Production capacity of 1 mm particles 80 kg / h; Dimensions of starting material -100 \ 60 mm, The upper limit of grinding - 2 mm, lower -1 mm; Effective chamber capacity - 1350 cm3 | 2014 | 2015 | Crusher used for crushing solid materials, stone, concrete, mountain rocks | Scientific, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Mixer machine tool model 022 Usman plant of foundry equipment, Energy industry | Voltage 380 V. Frequency 50 Hz. Rotation speed - 1390 rpm, volume 70 l | 2014 | 2015 | The mixer is used for intensive mixing of liquid and solid mixtures | Scientific, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Laboratory instrument "Vibro-mill", Scientific-technical company „Etalon“ | Grinding chamber 2 pcs., One chamber capacity 1 l, axial grinding size max. - 20 mm, min - 2 mm., Power -55 kV, voltage -380 V. Frequency -50 Hz. | 2015 | 2016 | Vibro mill is used for grinding solid particles | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Thermostat STI BASIC, POL-EKO Aparatura SP.J | Voltage 230 V, 50 Hz, Power 160 V. | 2016 | 2017 | Thermostat STI BASIC The purpose is drying of substances at a constant temperature, or thermostating of bio-soils | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Drying cabinet SNOL 60/300, SNOL | Power - 2 kW. Voltage-230V, frequency 50/50 Hz. Temperature - 50-3000C, change of air circulation in the chamber -500C-3 times / h, 1000C-8 times / h, 2000C- 12 times / h, 3000C - 16 times / h | 2019 | 2019 | High class Digital thermoregulatory dryer is used as an electric oven | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| High pressure autoclave with magnetic lock Toption GSH-3L, TOPTION GROUOCo, LIMITED, China | Maximum pressure -1.0 mg.pa, max. Temperature - 2500C, pressure - 0.8 mg. pa, working temperature -2000 C Nominal volume 3 l. | 2016 | 2017 | Autoclave is used to obtain chemically pure salts at high pressures and temperatures | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| High pressure autoclave with magnetic lock Toption GSH-5L, TOPTION GROUOCo, LIMITED, China | Maximum pressure -1.0 mg.pa, max. Temperature - 2500C, pressure - 0.8 mg. pa, working temperature -2000 C Nominal volume 5 l. | 2020 | 2021 | Autoclave is used to obtain chemically pure salts at high pressures and temperatures | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Controller D79350 Sexau, Plating Electronic GmbH, Germani | Voltage 1˜ 230V +/- 10%, 50-60 Hz, current-6.5 A, rectified voltage -0-12V, Power 0-60, Display resolution 10 mv / 100mv | 2017 | 2017 | The rectifier is used to conduct alternating current for electrochemical processes | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Power rectifier type Power Station pe 1028, Electronic GmbH | Voltage 1˜ 230V +/- 10%, 50-60 Hz, current-6.5 A, rectified voltage -0-15V, Power 0-30, Display resolution 10 mv / 100 | 2015 | 2015 | The rectifier is used to conduct alternating current for electrochemical processes | Scientific, Educational, industrial | Active use, working condition | G. Agladze Laboratory of Electrochemistry and Electrometallurgy
| Thermal analyzer NETSCH STA 25 Regulus, NETSCH | The temperature range of measurement is 25-11000 C. Heating rate 0.5-200 C. Research can be carried out both under normal conditions and in the required gaseous medium and vacuum. Samples for research are placed in aluminum and ceramic containers as needed. The maximum weight of the test sample is 0.5 g. The analysis process is automated and controlled by a computer. The results of the analysis are processed using the Proteus program. Research results are stored in memory. | 2017 | 2017 | Thermal analysis | Used for scientific educational purposes | Active use | N. Landia Laboratory of Thermochemistry
| Calorimeter HT-1500, French company- Seteram | Device that allows us to study a substance in a wide range of temperatures, from a room temperature to -1500K in a controlled atmosphere (vacuum or inert gas).The instrument consists of: High temperature graphite cylindrical furnace and graphite screen, thermal stacker - "Kalorimeter" regulator for obtaining different pressures, temperature regulator, amplifier, device – recorder. The calorimeter is placed inside the furnace in a corundum muffle, it consists of two cells containing two crucibles surrounded by 18 pairs (pt/Rh-pt/Rh10%) of thermobattery. The programmer-regulator, in turn, consists of a thermocouple "probe" (Pt-Pt/Rh 10%), which is placed in the middle of the muffle and regulates the furnace temperature. The regulator can maintain a constant temperature in the calorimeter (isothermal mode) or scan the temperature at different speeds, both during heating and cooling, the stability of the temperature value is equal to ± 0.001 C0. “Setaram's” calorimeter HT-1500 can be used to regulate and control the temperature, which in turn gives us the opportunity to conduct research in different modes. |  | 1975 | The isothermal mode is mainly useful in the study of interaction processes such as determining the heat of solubility of substances in each other, as well as determining the enthalpy, heat capacity and entropy of solids by the throwing method. The temperature scan mode at a speed of (20-2000) is used to determine the heat of structural transformation of substances (heat of ferromagnetic transformation, heat of structural transformation of substances, determination of enthalpy and entropy of order-disorder transformation), as well as the scan mode allows us to study the influence of heat treatment on the results of thermal transformations of substances. | Scientific | Active use | N. Landia Laboratory of Thermochemistry
| Calorimeter - DSC-111 products "Setaram". | The calorimeter is used to study the thermodynamic parameters of compounds or general physical and chemical systems under conditions of temperature, pressure and time changes. Operating temperature range 160-900K. | 1975 | 1982 | Used to obtain the following data: Temperature functions of heat capacity, thermodynamic characteristics of phase transformations (enthalpy, entropy, temperature regions of transition). | Scientific | Currently depreciated | N. Landia Laboratory of Thermochemistry
| Staimless distiller, China | 5 litre/hour; 220 V 4.5 kw | 2019 | 2020 | Distillation of water | Scientific | Active use | N. Landia Laboratory of Thermochemistry
| Universal Polarograph PU-1, Measuring Instruments Plant, Gomel, Belarus | Range of concentrations for determination of cadmium 10-3 ÷ 10-8. Range of sample concentration in the inverse preliminary accumulation mode 10-10 mol/L. Sensitivity of the polarograph during the direct current polarography is 80 mV/nA. Sensitivity of the polarograph during the alternating current polarography is 8 mV/nA. Alternating current power supply (220 ± 22) / (50 ± 0.5) W / Hz. Initial polarization voltage range -3 ÷ +3 W. Time interval measurement range 0,1 ÷ 990 sec. Power consumption not more than 60 watts. Overall dimensions of the measuring block 490×400×215 mm. Weight of the measuring block not more than 20 kg. | 1977 | 1978 | Intended for qualitative and quantitative analysis of compounds using voltammetry, and for another electrochemical research. For determining the impurities in metals, alloys, semiconductors, and chemical reagents. For controlling the purity of water, air, food, and medications. For conducting biochemical research. For studying electrodes, adsorption, and oxidation-reduction processes in chemistry of complex compounds. | Scientific, Educational, industrial | Active use, working condition | Jondo Japharidze elechtrochemistry of non-aqueous solutions laboratory
| Universal Polarograph OH-105, Promartel of electrical equipment "Redalkic" Budapest, Hungary | 1. Indicator electrode potential range Utot : +2÷ -6 V. 2. Polarizing voltage range Upol (mV/mm). 0 ÷ -0,5 V ±2 mV 1,25 mV/ mm. 0 ÷ -1,0 V ±2 mV 2,5 mV/ mm. 0 ÷ -2,0 V ±4 mV 5,0 mV/ mm. 0 ÷ -4,0 V ±8 mV 10,0 mV/ mm. 3. Starting voltage setting range Ust within +2.0 ÷ -2.0 V in steps of 0.2 V accuracy ±0.2% or 2 mV, whichever is greater. 4. Maximum cell voltage Umax: I = 0mA 50V min, I = 1mA 40V min, I = 5mA 2V min. 5. Amplitude of superimposed alternating voltage UAC eff: 2mV ±5%; 5mV ±5%; 10mV ±3%; 15mV ±3%; 20mV ±3%; 30mV ±3%. 6. Frequency of superimposed alternating voltage fpol: 60±1Hz. 7. Sensitivity when measuring direct current and alternating current 60 Hz within 3x10-10 ÷ 10.5 x 10-6 A / 0 can be adjusted in 108 steps with an accuracy of ± 2%. Reproducibility ± 0.5%. 8. DC voltage measurement sensitivity within 0.3 mV ÷ 3.5 V can be set in 96 steps with an accuracy of ± - %. 9. Permissible maximum measurable voltage: Uin max 350V. 10. Diffusion current compensation range:Idiff 0 ± 10-3 A. 11. Input impedance when measuring DC voltage: Rk 1 MΩ. 12. Capacitor current compensation range depending on the polarizing voltage: Ic/Upol 0 ÷ 5. 10-7 А/V. 13. Mains frequency 50Hz. 14. Permissible mains voltage fluctuation: ± 10%. 15. Standard temperature limits: ± 18 ÷ +280 C. 16. Operating temperature limits: ± 5 ÷ +350 С. 17. Overall dimensions: 540 x 420 x 175 mm. 18. Weight approx. 10 kg. 19. Protection class against electric shock 1. 20. IP-20 protection. | 1982 | 1984 | Polarography is an electrochemical test method in which, by applying an increasing DC voltage to the system under test, the strength of the current flowing is recorded. Electrodes are inserted into the electrochemical cell with the system under test. One of which, as a rule, is a mercury drop electrode with a small surface area, which acts as a cathode. The second electrode is bottom mercury with a large surface or a platinum electrode acting as an anode and a reference electrode, such as a calomel electrode. In order for the diffusion process to be independent of the process of ion movement caused by an electric force field, a conductive salt solution, a background electrolyte, is added to the polarographed sample. Classical polarography is based on the registration of the diffusion current, which increases with increasing supply voltage and reaches its limiting value. As the concentration of the test substance increases, the current limit value increases, which is the basis of the calibration curve in quantitative analysis. The value of the half-wave potential serves as a qualitative characteristic of the test substance. Under normal conditions, polarographically non-reducible substances under the action of certain catalysts cause the appearance of a step and, on the basis of the force of the so-called catalytic current, that appears in this case, the concentration of the catalyst can determined. Various inorganic ions, primarily metal ions, can be easily identified by polarographic method, but by including a dropping mercury electrode at the anode, anions can be analyzed. Thus, organic compounds dissociated into ions and non-ionic organic compounds can be determined polarographically if their molecule contains a radical that can be reduced or oxidized on a mercury drop electrode. Polarography can be applied in the analysis of weak solutions, microanalysis and trace analysis. The use of a sinusoidal AC voltage of small amplitude superimposed on a DC voltage (AC polarography) has a number of advantages. In the optimum case, components with a half-wave difference of approx. 50 mV can be determined next to each other. The relative amount of the determined components influences the measurement accuracy to a lesser extent. | Scientific, Educational, industrial | Active use, working condition | Jondo Japharidze elechtrochemistry of non-aqueous solutions laboratory
| Universal ionomer ЭВ-74, Measuring Instruments Plant, Gomel, Belarus | 1. The measurement limits of the pX value by the transducer are from minus 1 to plus 19 pX with ranges: -1-4 pX, 4-9 pX, 9-14 pX, 14-19 pX and a wide range: -1-19 pX. The limits of measurement of the pX value by the ion meter and the type of ion are determined by the type of measuring electrode used in the set of the ion meter. 2. The measurement limits of the Eh value (emf) by the transducer are from minus 100 to plus 1900mV and from plus 100 to minus 1900 mV with ranges: -100 ÷ 400 mV or 100 ÷ -400 mV; 400 ÷ 900 mV or -400 ÷ -900 mV; 900 ÷ 1400 mV or -900 ÷ -1400 mV; 1400 ÷ 1900 mV or -1400 ÷ -1900 mV and wide range: -100 ÷ 1900 mV or 100 ÷ -1900 mV. 3. The transducer in pX measurement mode provides operation with the following electrode systems: a. systems with normalized values of the coordinate of the isopotential point Ei and pXi (for example, systems with glass measuring electrodes for measuring pH, pNa, pK, etc.) with a characteristic determined by the equation: E=Eu+[S20+α(t-20)](рХ-рХи), (1) where E is emf of the electrode system, mV; S20 is the steepness of the characteristics of the electrode system at a temperature of 20ºC, mV/pX; t is the temperature of the solution, °С; Eu and pXi are the coordinates of the isopotential point of the electrode system, mV and pX, respectively; α is the temperature coefficient of steepness, equal to 0.1984 for monovalent ions, and 0.992 for divalent ions. b. systems that do not have normalized coordinates of the isopotential point (for example, systems with measuring membrane electrodes for measuring pJ, pCl, pCN, etc.), with a characteristic determined by the equation: E = E0 + S pX, (2) where E0 is emf of the electrode system at рХ=0, mV; S is the steepness of the characteristics of the electrode system, mV/pX. 4. The transducer in the millivoltmeter mode provides operation with electrode systems for measuring Еh, as well as (according to calibration curves) with electrode systems for measuring рХ. 5. The converter provides operation with systems having electrical resistance of the measuring electrode circuit - no more than 1000 MΩ, auxiliary electrode circuit - no more than 20 kOhm. 6. The limit of the permissible error of the ionomer when measuring pH using buffer solutions of the 2nd category is ±0.05 pH. 7. The transducer registration settling time (t set) does not exceed the value determined by the formula: t = (5 + 10R) s, where R is the numerical value of the internal resistance of the emf source, GΩ. 8. The stability of the transducer for 8 hours of continuous operation is not worse than ±1 mV. Output voltage fluctuations, no more than ±0.5 mV. 9. The output voltages of the converter when the pointer of the indicating device is set to the end of the scale are equal to: a) on sockets "2V" - 2.00 ± 0.05V in the range of 20 pX (2000 mV) with a reading of 14 pX; b) on sockets "20mV" - 20.0 ± 0.2 mV - on all ranges. 10. The power consumed by the ionomer does not exceed 50 VA, including 15 VA consumed by the converter. 11. Overall dimensions, mm, not more than: a) transducer 365×230×260; b) complete stands 260×260×500. 12. The weight of the ion meter is not more than 15 kg, including the converter, not more than 7 kg. | 1987 | 1991 | The universal ionomer EV-74 is designed to determine, together with ion-selective electrodes, the activity of mono- and divalent anions and cations (pX values) in aqueous solutions, as well as to measure redox potentials in the same solutions. The ionomer can also be used as a high-resistance millivoltmeter. When working with an automatic titration unit, the device can be used for mass titration of the same type. The EV-74 ionometer can be used to measure both by sampling and directly in laboratory facilities. The ionomer is designed for use in laboratories of research institutions and industrial enterprises. | Scientific, Educational | Active use, working condition | Jondo Japharidze elechtrochemistry of non-aqueous solutions laboratory
| Alternating current bridge Р5021, precision Electrical Appliance Plant, Ukraine, Kyiv. | 1. Power frequency area of the bridge measuring circuit depending on the measured parameters from 20 to 200 000 Hz. 2. The equivalent scheme for replacing the object of measurement is represented by three types of conductivity and capacitance connection: serial (connection of serial conductivity G1 and capacitance C); parallel (connection of parallel conductivity G2 and capacitance C); parallel-serial (a sequential combination of parallel-connected parallel conductivity G2 with capacitance C and sequential conductivity G1). 3. Bridge measurement ranges: by capacity (C) from 0.01 to 100 μF; by sequential conductivity G1 from 10-4 to 10 Ohm; by the parallel conductivity G2 from 10-4 to 1 Ohm. 4. The limit of the basic error in measurements according to the serial-parallel replacement scheme on all subbands should not exceed ± 15%. | 1963 | 1963 | The alternating current bridge P5021 is designed for electrochemical studies by measuring the impedance components of the electrochemical cell using serial, parallel and serial-parallel replacement circuits with simultaneous polarization of the electrochemical cell with a direct current voltage from a polarizing voltage source or from a potentiostat (not included in the bridge set) relative to the comparison electrode. The bridge is designed for operation in laboratory conditions at ambient temperature from + 10 to + 35 0 С and relative humidity up to 80% (at temperature + 20 0 С) in an atmosphere that does not contain acid vapors and other substances that cause corrosion. | 1. Scientific - used to determine the activity (pX values) of organic substances and drug forms in aqueous and non-aqueous solutions, buffers and complex compounds. 2. Scientific - used to study the electrochemical properties of organic substances and medical preparations, in particular the adsorption capacity in tandem with an alternating current probe, as a high-impedance millivoltmeter. It is also used to study the adsorption of inorganic substances and dosage forms in aqueous and non-aqueous solutions, in buffer and complex solutions. 3. Educational - for masters and students to acquire the skills of working with a pH-meter in chemical solutions and biological materials during the study of organic substances and pharmaceutical forms. 4. Educational - for masters and students, acquiring and deepening the skills of working in tandem of alternating current bridge with the pH-meter in studying the adsorption capacity of substances. | Active use, working condition | Jondo Japharidze elechtrochemistry of non-aqueous solutions laboratory
| Chromatograph GeC - 2014 , Firm “Shimadzu”, Japan | With heat conduction detector, microcomputer column ОD 1/16, carrier gas - helium, software support. Analyzes: CO, N2, O2. | 2021 | 2021 | Study of catalyst activity in the CO oxidation reaction | Scientific | Working condition | Catalysis Laboratory
| Mufel laboratory oven MP - 2 YM, Lithuania, the city of Utena | Power 2600 watts, operating temperature 10000 C | 1979 | 1980 | Temperature formation of catalysts | Scientific | Working condition | Catalysis Laboratory
| Chromatograph „Gazochrom” 3101, Factory “Chromatograph” Moscow | Combustion Heat Detector, Analyzes: CO, H2, CH4, O2, Accuracy 10 -3 | 1977 | 1978 | Study of catalyst activity in the CO oxidation reaction | Scientific | Working condition | Catalysis Laboratory
| Electronic drying cabinet 2V-151, USSR, Odes | Working temperature | 1976 | 1976 | Formation of catalysts | Scientific | Working condition | Catalysis Laboratory
| Muffle furnace PEM-2/86, firm “PRODRIN” (Poland) | Dimensions 180mm x 390mm x 125mm; chamber pan area 7,02 dm2; chamber volume 8,77 dm3; current characteristics, one-phase, 50 Hertz; voltage, 220 V; nominal power 3,6 kWt | 1985 | 1985 | The furnace is mainly intended for thermal processing (roasting, reduction, oxidation, sintering) of different materials and solid mixtures | Scientific, as well as industrial (in factory laboratories) | Active use | L. Japaridze Laboratory of Electrochemical Power Sources and Electrode Material
| Vibration mill, Teplogorsk Hydroequipment Plant, Russia | Initial particles size, no more than 3mm; vibrator’s number of revolutions 1400 rpm; steel pads number – 4 pieces; electric motor capacity, 0,6 kWt; mill weight 129 kg | 1979 | 1979 | Ore and other dry material grinding for chemical analysis | Scientific, industrial | Active use | L. Japaridze Laboratory of Electrochemical Power Sources and Electrode Material
| Crusher ДЛЩ 80Х 150А, “Amursky metallist” (Russia, Blagoveshchensk) | Output, no less than 1100 kg per h; outlet port length, from 1 to 20 mm; engine strength, 1,5 kWt; maximal size of charging material pieces, no more than 70 mm;crusher weight 160 kg. | 1987 | 1987 | The crusher is used for fine crushing of ores and other solid materials under laboratory conditions | Scientific, as well as industrial (in factory laboratories) | Active use | L. Japaridze Laboratory of Electrochemical Power Sources and Electrode Material
| Rectifying device , Electrotechnical plant, Russia | BAKPM-3200 – 12У4; model BAKPM-3200 – 12У4; circuit voltage 380 V; rectified voltage 12 V; rectified current 3200A; efficiency coefficient 83%; weight 760 kg | 1985 | 1985 | Conducting of electrolysis, receipt of galvanic coatings | Scientific, industrial | Working condition | L. Japaridze Laboratory of Electrochemical Power Sources and Electrode Material
| Rectifying device, Tambov electrotechnical plant | Model TEP-400/12; circuit voltage 380 V; number of phases – 3; rectified voltage 12 V; rectified current 400 A; efficiency coefficient ~85% | 1989 | 1989 | Conducting of electrolysis, receipt of galvanic coatings | Scientific, industrial | Active use | L. Japaridze Laboratory of Electrochemical Power Sources and Electrode Material
| Laboratory ball mill „МБЛ“, Plant Reductor, Russia | Number of chambers – 2; inner diameter of chamber 500 mm; length 280 mm; output (according to clinker) 5 kg per h; drum rotation velocity 48 rpm; electric power 1,1 kWt; weight (without reduction gear and engine) 438 kgdimensions 1305mm x 868m x 731mm; | 1986 | 1986 | Grinding of different materials under laboratory conditions according to dry and wet method (clinker, coke, different ores) | Scientific, as well as industrial (determination of degree of materials grinding) | Active use | L. Japaridze Laboratory of Electrochemical Power Sources and Electrode Material
| X-ray device DRON – 3M, Leningrad Optical and Mechanical Association | Range of rotation of the detection: (0÷ -1000 და 0÷ +1670); Detector displacement rate - 0.01÷1.25; Duration of continuous work no more than 12 hours; The energy consumed by the device is no more than 6 kilowatts; Radiation power of the X-ray source - 3 kilowatts; High voltage range - (2 - 60) k. volts; Anode current range - (2-60) milliamperes. | 1987 | 1987 | The device is used for the study of crystalline substances (solid powders), for the identification of individual substances included in them, for determining the phase composition of substances; for determining the crystal modification and structure of substances. The roentgenogram obtained on the device also allows to evaluate the degree of dispersion of crystalline substances. | Scientific research | Needs repair (X-ray tube needs to be replaced) | Laboratory of Physico-chemical analysis
| Differential scanning microcalorimeter DCM-2M, Special Design Bureau for Biological Instrumentation, Pushkino | Working temperature range of the tool - 123÷773 K; Heating rate - 0.5, 1, 2, 4, 8, 16, 32, 64 K/min; Alternating current from the network 220±10 V, frequency - 50±1 hertz; Consumed power - 200 VA; Overall dimensions - 745 X 230 X 320 mm; Weight - 30 kg. | 1978 | 1986 | The purpose of the device is to measure the heat capacity and enthalpy of the research samples, to determine the melting and crystallization temperatures. | Scientific research | Needs repair (amortized | Laboratory of Physico-chemical analysis
| Derivatograph Q-1500D; Optical Instrument Factory, Budapest, Hungary (MOM) | Measurement interval 0÷10 g.; Sensitivity 50, 1090, 200, 500, 100, 2000 mg/250 mm; Heating temperature interval 125, 250, 500, 1000, 15000 C; The maximum oven temperature is 10000C; Internal dimensions of the oven: 70X120 mm; Working atmosphere - air, inert gas; Supply voltage: 220±10 V, 50 Hz; Power consumption of the tool 4000 VA; | 1982 | 1986 | The instrument is used to measure the temperature (T) of the test sample, as well as to measure the mass change (TG), rate of mass change (DTG) and value of heat capacity (DTA) as a function of time. | Scientific research | In working condition | Laboratory of Physico-chemical analysis
| Atomic absorption spectrometer ANALYST 200; PerkinElmer - USA | Power supply from the network 115-230±10 V, 50 Hz; Power consumption 300 VA; Delivered gases - air + acetylene; also works with air+N2O gas; Equipped with a pre-mixing burner. | 2013 | 2015 | Determination of quantitative concentration of various elements | Scientific research | Active use | Laboratory of Physico-chemical analysis
| Microwave digestion system MDC-6G; Shanghai Sineo Microwave Chemistry Technology (China) Co., LTD | Technical characteristics: power supply - 220-240 VAC 50/60Hz 8A; Microwave frequency - 2450 MHz; Power - 1800w; Temperature range 0 – 3000 C; Tool size - 450 x 515 x 510 mm; Tool weight - 40 kg. | 2017 | 2017 | Preparation of solutions for soluble substances, as well as for sparingly soluble and insoluble substances. Facilitates the process of dissolving substances and the simultaneous preparation of several different solutions. | Scientific research | In working condition | Laboratory of Physico-chemical analysis
| Atomic Absorption Spectrophotometer C-115 #79486; Special construction bureau - В-2613 | Power supply of the device from the network 220±10 V, 50 hertz; The tool is equipped with a compressor for supplying compressed air (2 at.); Working power 0.28 kW; The mass of the tool is 140 kg. | 1985 | 2008 | The purpose of the device is to measure the concentration of chemical elements in liquid samples; for ore processing, ecological, agrochemical, metallurgical and other purposes. | Scientific and industrial | In working condition | Laboratory of Physico-chemical analysis
| X-ray fluorescence spectrometer EDX 3600 H, Skyray Instrument Co. LTD | Technical characteristics: measuring elements from Na to U; Sensitivity - 1 ppm - 99.99%; Simultaneous measurement of 24 types of elements; The thickness of the coating to be measured is up to 0.005 μm for the each layer; It is possible to measure more than 11 layers; X-ray tube voltage - from 5 kV to 50 kV. Current of the tube – from 50 ml. A to 1000 ml. A; Sample cell dimensions 320 mm x 100 mm, tool dimensions - 660 x 510 x 350 mm; | 2017 | 2017 | Elemental analysis of substances and alloys, measurement of metal coating thickness, ROHS analysis | Scientific and industrial | In working condition | Laboratory of Physico-chemical analysis
| Infrared spectrometer FT-IR TENSOR II; BRUKER OPTIK GmBH | The instrument is equipped with a sample placement device - ATR, which uses a diamond single crystal; beam splitter material - KBr; Radiation source - GLOBAR (SiC); Window material - KBr; Detector - DIGS-HgCdTe – InSb, with KBr window; Measurement range - 400 cm-1÷4000 cm-1. | 2018 | 2018 | The device is intended for studying the structure, composition, structural units (fragments) of substances, both in solid state (crystalline and amorphous) and in solutions, by measuring the infrared spectrum. IR spectra also provide information about the state of molecules adsorbed on the surface of the substance and present in its volume. | Scientific and industrial | In working condition | Laboratory of Physico-chemical analysis
| Precision PH meter-ionometer Sanxin MP523; Shanghai San-Xin Instrumentation, Inc | Measurement range - (-1.999÷19.999) pH (mV); Temperature range - -100C ÷ 1100C; Power - DC9V/300mA; Tool size - 160 x 190 x 70 mm; Mass - 880 g. | 2019 | 2019 | The device is intended for determining the pH of solutions, for measuring oxidation-reduction potential. Also, in the presence of a suitable ion-selective electrode, to determine the concentration of ions. | Scientific and industrial | In working condition | Laboratory of Physico-chemical analysis
| Water purification system Labaqua HPLC; Biosan SIA, Riga, Latvia | Incoming current 200-240 V, 50-60 Hz, 130 watts; inlet water pressure - 0.5 - 5 bar, electrical conductivity - ≤ 1300 µS/cm; Connection of incoming water to the system - with ǿ 1/2'' NPTF; Dimensions of the tool - 320x560x620 mm, weight - 25 kg; Water collection volume - 30 l. | 2019 | 2019 | The purpose of the water purification system is to obtain clean water that meets ISO 3696 - I and II quality requirements and is used for the preparation of precise solutions in highly sensitive analytical techniques (atomic-absorption, spectrophotometric, ICP-MS and others); | Scientific and industrial | In working condition | Laboratory of Physico-chemical analysis
| Spectrophotometer, Zuzi 4201/50 | Power: AC 220V/50Hz; Dimensions - 480x360x160mm; Wave length - 200-1000 nm; Wavelength accuracy - ±2nm; Sample section - 4 cells of 10 mm; Lamps: tungsten and deuterium; Photometric range: -0.097/1.999A; 0/125% T. | 2008 | 2017 | The instrument can perform accurate qualitative and quantitative analysis in the 200-1000 nm range. | Scientific | Active use | R. Dogonadze Laboratory of Theoretical Investigations
| Trinocular metallurgical microscope with digital camera, Euromex ME 2665 | Power source 230V, 50/60Hz; eyepiece -10; Condensed lens Abbe (N.A.25, with adjustable lens); Halogen (6V/30W, Kohler lighting); Lighting - delivered and transmitted light; Magnification: 40, 100, 400. | 2015 | 2017 | A microscope is used to examine non-transparent and transparent materials (metal, plastic, minerals, asbestos, glass, wood, chemicals). | Scientific | The tool is in working condition | R. Dogonadze Laboratory of Theoretical Investigations
| Trinocular microscope, with digital camera, Am scope T690C-PL10m (HOLLAND) | Power source 230V, 50/60Hz; eyepiece -10; Halogen (6V/30W, Kohler lighting); Lighting - delivered and transmitted light; Magnification: 40, 100, 400, 1000. | 2011 | 2017 | A microscope is used to examine non-transparent and transparent materials | Scientific | Active use | R. Dogonadze Laboratory of Theoretical Investigations
| Constant current and voltage controller ТЕС-21-23, Research and production enterprise "ANALYTIC" Mikhailovgrad | Current range - 0.05-2.5 A; Voltage range - 0.1-30 V. | 1983 | 1985 | Obtaining constant current and voltage | Scientific | Active use | R. Dogonadze Laboratory of Theoretical Investigations
| Regulated power supply 3010 D , ARKSEN | Power supply voltage –– 220 V / 50Hz or 110 V / 50Hz ±10% ; Working condition – o C -10-40; Dimension - mm - 285 x 128 x 145; Output voltage - 0-30 V; output current - 0-10A. | 2017 | 2018 | Constant current source | Scientific | Active use | R. Dogonadze Laboratory of Theoretical Investigations
| Combustion Chembar, Entrepreneur Mamuka Chigogidze | The chamber is made of a 2 mm high quality stainless steel tube (length 640 mm, diameter 170 mm). There is a quartz glass porthole on the front of the pipe for viewing. A chimney (corrugated aluminum pipe) is attached to one end of the pipe. A metal stand for an ash pan is placed in the chamber, with a moving gas lighter (lighter - CHFS) attached to the stand. An electric heater (with nichrome coils) and a chrome-alumel thermocouple connected to a millivoltmeter are installed in the chamber. For thermal insulation, the pipe is wrapped with a thermal insulator. | 2017 | 2017 | The combustion chamber is used to study the flammability of the material and the speed of flame propagation on the surface of the material. | Scientific | Needs repair | R. Dogonadze Laboratory of Theoretical Investigations
| Equipment for charging fire extinguisher cylinders, Entrepreneur Mamuka Chigogidze | The device consists of a cylinder with nitrogen or CO2 (the manometer is attached to the cylinder reducer) and a holder, which is a structure that holds the fire extinguisher cylinder and provides a reliable, convenient and safe position for charging it. Powder is poured into the fire-extinguishing cylinder, the cylinder is capped, an armored hose is screwed into it, and the cylinder is fixed in the holder. The other end of the armored hose is screwed onto the CO2 cylinder. The valve of the fire extinguisher cylinder is opened with the handle and fixed. Then the CO2 cylinder is opened and the fire extinguisher cylinder is filled with CO2 until the arrow of the manometer on the reducer enters the green zone. | 2017 | 2017 | Extinguishers are used to charge fire extinguisher cylinders with fire extinguishing powders. | Scientific | Working condition | R. Dogonadze Laboratory of Theoretical Investigations
| Fire extinguishing capacity determination tool, Entrepreneur Mamuka Chigogidze | Tube made of 2 mm high quality stainless steel (length 600 mm and diameter 150 mm). The pipe is covered with an autodynamic sieve (sieve - 0.3 mm) with a mechanical stirrer controlled by a variable frequency generator (the generator consists of two TVK transformers -110 m (2 - 9 V)). | 2017 | 2017 | The fire-extinguishing ability determination tool is used to study the effectiveness of fire-extinguishing powders, the extinguishing mechanism and the optimal extinguishing conditions. | Scientific | Working condition | R. Dogonadze Laboratory of Theoretical Investigations
| Tumbling mill, Entrepreneur Mamuka Chigogidze | The mill consists of a drum, an electric motor, a reducer and a variator. The drum is a steel cylinder with a diameter of 500 mm and a length of 600 mm. For loading and unloading, the drum has a rectangular hole with a universal lid, on which a set of sieves of different sizes is attached, which ensures the sifting of powders at the moment of grinding. The purpose of the reducer and variator is to reduce the rotation frequency of the electric motor from 1500 rpm to 100-20 rpm. We put the grinding material in the drum, throw steel balls of different sizes (70-10 mm in diameter) weighing 100-130 g and turn on the machine. The duration of grinding depends on the dispersion of the obtained powder. | 2018 | 2018 | Tumbling mill is used for grinding hard and medium hard rocks (limestone, dolomite, sandstone, zeolite, diatomite and others). This mill, unlike standard mills, allows us to sift and sort the powders according to the required dispersion at the moment of grinding. | Scientific | Working condition | R. Dogonadze Laboratory of Theoretical Investigations